Optimization of the effective light attenuation length of YAP:Ce and LYSO:Ce crystals for a novel geometrical PET concept

I. Vilardia, A. Braemb, E. Chesib, F. Ciociaa, N. Colonnaa, F. Corsic, F. Cusannod, R. De Leoa, A. Dragonec, F. Garibaldid, C. Joramb, L. Lagambaa, S. Marronea, E. Nappia, J. Séguinotb, G. Taglientea, A. Valentinia, P. Weilhammerb, H. Zaidie

aPhysics Dept. and INFN Section of Bari, Via Orabona 4, Bari, Italy

bCERN PH-Department, CH-1211 Geneva, Switzerland

cDEE Politecnico di Bari – MICROLABEN s.r.l., Bari, Italy

dLaboratory of Physics, ISS, Viale Regina Elena 299, Rome, Italy

eDivision of Nuclear Medicine, Geneva University Hospital, CH-1211 Geneva, Switzerland

Abstract

The effective light attenuation length in thin bars of polished YAP:Ce and LYSO:Ce scintillators with lengths of the order of 10 cm has been studied for various wrappings and coatings of the crystal lateral surfaces. This physical parameter plays a key role in a novel 3D PET concept based on axial arrays of long scintillator bars read out at both ends by Hybrid Photodetectors (HPD) since it influences the spatial, energy, and time resolutions of such a device. In this paper we show that the effective light attenuation length of polished crystals can be reduced by wrapping their lateral surfaces with Teflon, or tuned to the desired value by depositing a coating of Cr or Au of well defined thickness. The studies have been carried out with YAP and LYSO long scintillator bars, read out by standard photomultiplier tubes. Even if the novel PET device will use different scintillators and HPD readout, the results described here prove the feasibility of an important aspect of the concept and provide hints on the potential capabilities of the device.

PACS: 29.40.Mc; 29.30.Kv; 87.58.Fg

Keywords: YAP:Ce; LYSO:Ce; PET; HPD.

# Introduction

The CIMA collaboration [1] has proposed a novel 3D PET geometrical concept [2] based on axially oriented arrays of long polished scintillator bars (e.g. 3.2×3.2×100 mm3), read out at the  two ends by Hybrid Photodetectors (HPD) [3]. A schematic view of the novel device is shown in Fig. 1. The unambiguous definition of the -ray interaction point in a real 3D geometry eliminates [4] the parallax error due to the unknown depth of interaction (DoI) of the -ray thus improving the spatial resolution, sensitivity and contrast of the PET performance in the field of molecular imaging. The concept is new compared to available PETs which are based [2] on radial crystal arrangements and block readout schemes (Anger logic) [5]. It is conceptually different from the phoswich technique [5,6] which also aims to measure the DoI. The 3D Axial PET concept provides higher efficiency, due to the absence of limitations imposed by the detector thickness in the radial direction, and to the possibility [4] to recover a fraction of ’s undergoing double interactions (first Compton and then photoelectric in a different crystal of the same array).

Scintillation light produced in the crystal propagates to the two bar ends by internal reflection from the crystal lateral surfaces. On its way to the photodetectors, part of the light is absorbed with a characteristic attenuation length. While the transverse coordinates (x, y) of the detected -ray are determined from the address (i.e. position) of the hit crystal, deduced by both photodetectors with a spatial resolution [2] x,y = 3.2 mm /12 = 0.92 mm, the axial (z) coordinate is derived with a precision z from the ratio of the photoelectron yield, N1 and N2, measured at the two ends of the long crystals:

z=12λeffln  N1N2+LC  ,      σz=λeff2Noexp  zλeff+exp  LC−zλeff12,                                      (1)

with LC being the length of the crystal. eff is the effective light attenuation length and is different from the bulk value bulk as it takes into account the real path of the photons, which is increased because of the multiple bounces: eff  0.8·bulk. The expression for z accounts for the photoelectron statistics, however it ignores the fluctuations of the path lengths w.r.t. the average value.  N0 is the number of photoelectrons for eff → . Its value depends both on the physical and optical properties of the chosen scintillator (incl. surface coating/wrapping) and the characteristics of the photodetector. It is implicitly assumed that the photoelectron yields N1 and N2 depend exponentially on the average path length of the scintillation photons belonging to one -ray:

N1=N02exp−zλeff  ,          N2=N02exp−(LC−z)λeff  ,                               (2)

giving the total number of photoelectrons Npe:

Npez=N1+N2.                                                            (3)

The statistical term of the energy and time resolutions can be expressed as:

σEE=ENFNpe  ,          σT=cNpe  ,                                                    (4)

where ENF is the excess noise factor [7] of the photodetector and c is a constant. From the previous equations it can be seen that while an increase of N0 improves all the resolutions, an increase of eff, although improving E/E and T, worsens z.

In order to achieve a competitive z, the crystal length LC needs to be limited to values below ~150 mm and eff has to be optimized depending on the chosen scintillator and crystal length. The two parameters LC and eff play obviously a key role in the proposed geometrical PET concept.

In this paper we present an experimental study to optimize the effective light attenuation length, eff, for a set of polished YAP:Ce scintillators of dimensions 3.2×3.2×100 mm3 produced by Preciosa Crytur Co at Turnov, Czech Republic, and for a few samples of polished LYSO:Ce of equal size, produced by Photonic Materials, Bellshill, Scotland. We have measured the eff parameter in polished and wrapped crystals by means of 511 keV -rays emitted from a 22Na source.

Techniques to modify the reflectivity of the crystal by wrapping, coating, and, mostly, roughing its surface have been described [8] in the literature. We show here that, by coating the lateral surfaces of the polished crystals with layers of evaporated Cu or Au of various thicknesses, it is possible to tune eff to achieve a value which leads to optimized spatial, energy, and time resolutions.

A simulation with the photo tracking code Litrani [9] has been used to extrapolate the present study to other crystal coatings. The code will be used to predict results for other crystal lengths and for other scintillators, such as LaBr3 and LSO. The latter ones, having both high photofraction and light yield, appear hence more suited for PET devices. Experimental studies with these crystals are in progress.

Studies reported in the literature about PET designs [8, 10-13] where the DoI information is obtained reading out both crystal ends [14], refer to conventional PET geometries with shorter scintillator bars (2-3 cm) and a surface finish different (raw) from ours, thus favouring diffuse reflection of scintillation light from the crystal lateral surfaces, rather than specular one.

# EXPERIMENTAL SET-UP AND  RESULTS

## Experimental apparatus

We have carried out the experimental studies with YAP and LYSO crystals of dimensions 3.2×3.2×100  mm3. Their refractive indices (at the emission wavelength maximum) are nLYSO = 1.82 (420 nm), and nYAP = 1.94 for YAP (370 nm). The thickness of the crystal has been chosen to obtain an excellent transverse spatial resolution and is matched to the segmentation of the HPDs that will be used in the final project [2, 3]. A set-up has been built (see Fig. 2), which allows measuring the effective attenuation lengths. The crystal under test is read out by two photodetectors and is placed on a linear stage remotely movable by a step motor.

For the present study two H3164-10 PMTs with bialkali photocathodes and borosilicate windows (nwin = 1.47) have been used as photodetectors. The PMTs are coupled to the scintillator with optical grease (BC-630 from Bicron, nBC-630 = 1.47). The HPDs of the final device will be equipped with a thin sapphire window (n = 1.793 at 370 nm, n = 1.782 at 420 nm), thus leading to an almost perfect refractive index matching which minimizes transmission losses at the crystal-detector interface. In this latter case, an optical coupling oil, like the one commercialized by Cargille-Sacher Laboratories Inc [15],  with a refractive index higher than that of BC-630, will be used.

A point-like 22Na source sealed in plastic has been used. One of the two resulting 511 keV -rays is detected by the crystal that is placed in coincidence with a cylindrical BaF2 scintillator (2 cm diameter, 6 cm length) detecting the second 511 keV -ray. The BaF2 detector and the -source are placed at a fixed position outside of the linear stage. The radioactive source is collimated with a 4 mm lead shield block with a rectangular opening. It is positioned 45 cm from the BaF2 detector and 5 cm from the face of the tested YAP or LYSO crystal in order to illuminate a 2 mm wide slide of its lateral surface. Conventional NIM electronics has been used. The data acquisition has been restricted to coincident timing signals from the crystal-left, crystal-right and BaF2 PMTs. Timing signals have been formed by employing constant fraction discriminators (CFD). A scheme of the electronics used is included in Fig. 2.

An example of the obtained spectra (sum of the left and right signals) for polished YAP and LYSO crystals with 511 keV -rays impinging at the crystal centre (z = 5 cm) is shown in Fig. 3. The LYSO scintillator has a higher photofraction and produces more photoelectrons than YAP, as a consequence of a higher photon yield [2] (27 ph/keV for LYSO, 18 ph/keV for YAP), of a lower refractive index, and, as will be shown later, of a longer light attenuation length. But, in spite of the higher photoelectron number, the energy resolution of LYSO is worse than that of YAP, most probably due to a higher intrinsic resolution [2, 16]. The E/E (FWHM) values are: 10.8% for YAP and 14.6% for LYSO. This last value is comparable with those quoted in literature: 13-15% [10] and 14-18% [8], both obtained for 20 mm long LSO raw bars. The YAP energy resolution is better than the 14% value quoted in [16] for 30 mm long bars.

B.   Measurement of the effective light attenuation length eff

The signal amplitudes measured at left and right ends of the crystal bar are highly correlated. Fig. 4 illustrates this for a polished YAP scintillator exposed to 511 keV -rays impinging in the middle (z = 5 cm) of the crystal bar. The central spot in the figure corresponds to the photopeak. In Fig. 5 we plot Q, the charge of the photopeak (in ADC counts), as obtained from the left ADC spectrum, scanning polished YAP and LYSO crystals in the z range from 1 to 9 cm, in steps of 1 cm. The linear behaviour in log scale means that one exponential (Qo · exp[-z/eff]) is sufficient to describe the PMT pulse height as a function of the distance z of the -source from the PMT.

The average effective attenuation length evaluated on a set of sixteen polished YAP crystals results to be eff  = 20.80.4 cm. This value is higher than those reported in the literature [2, 16] and, as we will discuss below, would lead to a worse resolution of our z-reconstruction method. For a set of three polished LYSO crystals an even higher value, eff  =  42.00.9 cm, has been obtained. Also this value is higher than those (2-4 cm) that can be deduced from the data in [8, 10-13] that however refer to LSO scintillators not mechanically polished.

In Fig. 6 we report the results of the same z scans for YAP crystal bars with the lateral surfaces coated or wrapped. The experimental data show that also the behaviour of coated or wrapped crystals is well reproduced by one exponential. The resulting light attenuation lengths eff are reported in Table I. The investigated wrappings/coatings result in eff values in the range 3.9 to 11.9 cm, compared to 20.8 cm for the polished uncoated YAP crystal. Although each of the wrappings/coatings can effectively reduce eff, only the Teflon wrapping is able to practically halve it and to maintain, at the same time, a high light yield.

The metallic evaporation method has the advantage to allow tuning eff to a desired value by changing the coating thickness. This is shown in Fig. 7 where the experimental eff values obtained for Cr coated and polished YAP crystals are plotted and compared with the prediction of a photo tracking code [9], run with the following values: 24 cm for the YAP bulk attenuation length bulk, 0.109 nm for the Cr absorption length abs(370 nm), and 1.87 for the Cr refractive index nCr(370 nm).

The effectiveness of Teflon in reducing the photon attenuation length has also been proved for LYSO crystals. The effects of other wrappings/coatings have not been tested on LYSO because currently only few crystal samples are available. Presumably the effects are the same as for YAP.

We remark that the extrapolations of the data in Fig. 6 to z = 0 (which correspond to N0/2) show that the light collection efficiency is strongly dependent on the wrapping/coating. Simulations are ongoing in order to reproduce this reduced detected light yield.

C. Reconstruction of the z coordinate

The reconstructed z coordinate of the -ray interaction point in the crystal is derived from the ratio of signals at the left and right bar ends using Eq. 1. Fig. 8 shows the mean value of the reconstructed z coordinate for photopeak events, plotted as a function of the true z coordinate. These measurements result from averaging over many events, such that the statistical error of the mean value becomes negligible (smaller than the marker size in the figure). The upper panel of the figure is for a polished uncoated YAP and the lower one for a LYSO crystal. The data are very well described by a linear fit, proving the validity of the applied reconstruction method. A small deviation from linearity is observed towards the bar ends. Similar results have been obtained for coated and wrapped crystals. The non-linearity at the ends is slightly enhanced for the shorter eff values. The small discrepancy on the z-reconstruction at the bar end can be explained by the small increase of the solid angle in collecting the light.

The distributions of the reconstructed z points for individual 511 keV -rays are shown in Fig. 9 for a Teflon wrapped YAP crystal  for three different interaction depths (z = 2, 5, and 8 cm). These have a gaussian shape whose standard deviations, giving the z resolution in z-reconstruction, are shown in Fig. 10, both for YAP (left panels) and LYSO (right panels) crystals, in polished (upper panels) and Teflon wrapped (lower panels) condition. Values of z = 0.8 and 1.45 cm are found for YAP and LYSO, respectively, essentially constant along the crystal length for polished crystals. The two lower plots show a significantly improved resolution when the crystals are wrapped with Teflon tape: z = 0.6 and 0.95 cm. However the resolution slightly degrades towards the ends of the crystals. The measured z values in the crystal centre (z = 5 cm) are reported in Table I.

Smaller z values have been reported in [8, 10-12] for shorter LSO crystals. The best resolutions (z [FWHM] = 3-4 mm, [8, 10]) have been obtained for 2 cm long crystals with unpolished lateral surfaces, left “as cut by the saw”. A worse resolution (z [FWHM] = 10 mm) was obtained in [8] for a polished Teflon-wrapped 2 cm long LSO crystal.

The exponential law (Eq. 2) in determining the axial coordinate z (Eq. 1) provides higher precision than the linear approximation (z = Lc*N1/[N1+N2]) used in [8, 10-13] and based on the assumption N1(z)+N2(z)= const.  This assumption is not fulfilled by the exponential trend of all the experimental points shown in Fig. 6.

D.    Precision of z-reconstruction and energy measurement

To achieve optimum performance both for energy and z-reconstruction resolutions, the crystal has to provide high light yield (~N0) and adequate short effective light attenuation length (eff). These two requirements are to a certain extent contradictory and therefore a good compromise is mandatory.  The energy resolution E/E has been derived from the standard deviation of the photopeak of the sum spectra QR+QL, with a 22Na source positioned at the centre of the crystal (z = 5 cm). The plots in the upper part of Fig.s 11 and 12 show the energy resolution of YAP and LYSO crystals, respectively, as a function of the eff values obtained with the different coatings/wrappings. As expected, the energy resolution degrades with decreasing eff, while the z-resolution, measured at z = 5 cm (see the two central panels of Fig.s 11 and 12), shows the opposite behaviour, it improves with decreasing eff. The resolution in the time difference of the left and right PMT signals, measured at z = 5 cm, is displayed in the lower panels. Its behaviour is similar to that of the energy resolution. The interpretation of these plots indicates that the Teflon wrapping represents a good compromise for 10 cm long YAP crystals.  The experimental resolution values are summarized in Table I.

To check the internal consistency of our measurements, we derive the number of photoelectrons from the measured energy resolutions. An intrinsic resolution (E/E)intr of 1.4% and 5% has been assumed [2, 16] for YAP and LYSO, respectively, at 511 keV. The intrinsic resolution is unfolded, and the excess noise factor of the PMT (ENF = 1.4, [7])  is corrected for:

NpeLc2=ENF(σEE)2−(σEE)intr.2   ,                                                       (6)

N0=Npe(Lc2)⋅eLc2λeff.                                                                (7)

The obtained N0 values, listed in Table II, together with the measured values of eff, allow us to calculate the expected spatial resolution according to Eq. 1. We multiply the calculated z values by a factor of 1.13 to account for path-length fluctuations [2] which are not included in Eq. 1. The calculated values together with the measured ones are also summarized in Table II. We find good agreement on the level of 10%, except for 1 nm Cr coating on YAP, where the difference is about 20%. The agreement proves the validity of our approach.

# III     Conclusions

This paper is a status report on the various techniques we are exploring for the HPD-PET concept. It deals with the characterization of 10 cm long polished YAP and LYSO scintillators with 511 keV -rays.

The ratio of the N0 values of polished YAP and LYSO (1.36 from Fig. 3, 1.26 from Tab. II) is in good agreement with the ratio of their light yields quoted in the literature (1.5). Scans along the crystal bars revealed a fairly good exponential behaviour of the signal with the z coordinate. This behaviour is observed also after wrapping or coating the crystal lateral surfaces. This allows to reconstruct the z coordinate from the ratio of the signals measured at the two bar ends. This last result validates an important aspect of the 3D Axial PET concept.

The examined polished YAP and LYSO crystals were found to be significantly more transparent than those studied in earlier works. The polished LYSO crystals showed a very large effective optical attenuation length of about 40 cm, i.e. a bulk value of about 50 cm.  The polished YAP crystals showed a shorter, but still large, effective attenuation length of about 21 cm, i.e. a bulk value of about 25 cm. Wrapping or coating the crystal lateral surfaces allows decreasing the eff value, but also influences No. The two parameters affect z, E/E and t. The Teflon wrapping was found to be the best method to reduce eff and z while maintaining an acceptable N0 value and hence also good E/E and t resolutions. The Cr coating allows to tune eff to the desired value, but it decreases also N0 , and thus is less effective in improving z.

With Teflon wrapping E/E values of 4.7% and 7% have been obtained for YAP and LYSO crystals, respectively. The same wrapping gives z values of 5.4 and 9.3 mm for YAP and LYSO crystals, respectively. The worse LYSO values are most probably caused by a higher intrinsic (non-

statistical) term of this scintillator.

Both the YAP and LYSO z-resolutions are worse than values obtained in other PET design studies with a read out at both crystal ends performed with raw and shorter LSO scintillators. We plan to build the 3D Axial PET concept with custom designed HPDs, equipped with a thin sapphire entrance window, which leads to matched refractive indices. For Teflon wrapped YAP crystals, an improvement of 29% in the z values and of 40% in E/E are expected from simulation codes, which are due to a 50% increase of N0, a 13% decrease of eff, and to the ENF value of the HPD very near to 1.

Even if the best resolutions (E/E (FWHM) = 8%, z (FWHM) = 9.8 mm) for the HPD-PET concept are predicted for Teflon wrapped YAP crystals, the choice of the scintillator-wrapping to be used in the final project has not yet been made. Simulation calculations will be performed to extrapolate the presented results to other crystal lengths and to other scintillators. Experimental tests on LaBr3 and LYSO scintillators coupled to HPDs are under preparation. The effect of the variation of the Ce concentration on light yield [17] and optical attenuation length is another approach which we plan to study.

##### References

[1]      The CIMA collaboration: http://www.cima-collaboration.org/.

[2]      J. Seguinot et al., ”Novel Geometrical Concept of High Performance Brain PET Scanner –

Principle,  Design and Performance Estimates”, CERN preprint PH-EP/2004-050, submitted

to Nuovo Cimento.

[3]     C. Joram, “Hybrid Photodiodes”, Nucl. Phys. B (Proc. Supp.) 78 (1999) 407.

[4]      A. Braem et al., ”Novel design of a parallax free Compton enhanced PET scanner”,

NIM A 525 (2004)  268.

[5]      W.W. Moses,  “Trends in PET imaging”,  NIM A 471 (2001) 209-214.

[6]      G.F. Knoll, “Radiation Detection and Measurement”, John Wiley & Sons Inc., New York,

ISBN 0-471-07338-5, pag.344.

[7]      C. D’Ambrosio, H. Leutz, “Hybrid photon detectors”, NIM A 501 (2003) 463-498.

[8]      Y. Shao et al., “Dual APD array readout of LSO crystals: optimization of crystal surface

treatment”, IEEE Trans. Nucl. Sci. 49 (2002) 649.

[9]      F.X. Gentit, “The Litrani code”, http://gentit.home.cern.ch/gentit/.

[10]    P.A. Dokhale et al., “Performance measurements of a depth-encoding PET detector

module based on position-sensitive avalanche photodiode read-out”,  Phys. Med. Biol. 49

(2004) 4293.

[11]    J.S. Huber et al., “A LSO Scintillator Array for a PET Detector Module with Depth

of Interaction Measurement”, IEEE Trans.Nucl.Sci. 48 (2001) 684 ; G.C. Wang et al.,

“Calibration of a PEM Detector With Depth of Interaction Measurement”, Trans. Nucl.

Sci. 51 (2004) 775.

[12]   E. Gramsch et al., “Measurement of the Depth of Interaction of an LSO Scintillator

Using a Planar Process APD”, IEEE Trans. Nucl. Sci. 50 (2003) 307.

[13]   W.W. Moses and S.E. Derenzo, “Design studies for a PET detector module using a PIN

photodiode to measure depth of interaction”, IEEE Trans. Nucl. Sci. 41 (1994) 1441.

[14]    K. Shimizu et al., “Development of 3-D detector system for positron CT”, IEEE

Trans. Nucl. Sci. 35 (1988) 717.

[15]    http://www.2spi.com/catalog/ltmic/cargille-standard.shtml;

http://www.2spi.com/catalog/ltmic/cargille-liquid.html.

[16]    A. Del Guerra, et al.,  “Measurement of absolute light yield and determination of a lower

limit for the light attenuation length for YAP:Ce crystal”, IEEE Trans. Nucl. Sci. 44 (1997)

2415.

[17]     J.A. Mares, et al., “Scintillation and spectroscopic properties of Ce3+ -doped YAlO3 and

Lux(RE)1-xAlO3 (RE=Y3+ and Gd3+) scintillators”, NIM A 498 (2003) 312.

Table captions

Table I: Measured eff values and spatial, energy and time resolutions for 10 cm long YAP and LYSO crystals with different coatings. All measurements have been performed with -rays of 511 keV in the centre of the crystal. All resolutions have been evaluated from the standard deviation of the photopeak of the sum spectra N1+N2.

Table II: Comparison of expected and measured z-resolutions for YAP and LYSO crystals with various coatings for scintillations of 511 keV -rays in the centre of the crystal.

|  | eff (cm) | E/E (%) | z (mm) | t (ps) |
| --- | --- | --- | --- | --- |
| YAP Polished | 20.8 ± 0.4 | 4.6 ± 0.4 | 8.2 ± 0.3 | 440 ± 15 |
| YAP Cr (1 nm) | 11.9 ± 0.3 | 5.2 ± 0.3 | 7.1 ± 0.4 | 480 ± 15 |
| YAP White Painted | 10.7 ± 0.2 | 6.2 ± 0.3 | 6.5 ± 0.2 | 550 ± 20 |
| YAP Tefl. wrap. | 10.5 ± 0.3 | 4.7 ± 0.2 | 5.4 ± 0.3 | 470 ± 20 |
| YAP Black Painted | 5.4 ± 0.2 | 10.6 ± 0.7 | 6.1 ± 0.3 | 775 ± 30 |
| YAP Au (1.5 nm) | 5.2 ± 0.2 | 7.7 ± 0.8 | 4.8 ± 0.4 | 640 ± 35 |
| YAP Cr (3 nm) | 3.9 ± 0.2 | 13.4 ± 1.5 | 5.3 ± 0.2 | 925 ± 30 |
| LYSO Polished | 42.0 ± 0.9 | 6.2 ± 0.2 | 14.4 ± 0.4 | 420 ± 10 |
| LYSO Tefl. wrap. | 20.0 ± 0.5 | 7.0 ± 0.2 | 9.3 ± 0.3 | 490 ± 10 |

Table I

|  | eff (cm) | E/E (%) | No | z (mm) expected | z (mm) measured |
| --- | --- | --- | --- | --- | --- |
| YAP polished | 20.8 | 4.6 | 927 | 8.7 | 8.2 |
| YAP Tefl. wrap. | 10.5 | 4.7 | 1120 | 4.5 | 5.4 |
| YAP Cr (1nm) | 11.9 | 5.2 | 850 | 5.65 | 7.1 |
| YAP Cr (3nm) | 3.9 | 13.4 | 284 | 5.0 | 5.3 |
| LYSO polished | 42.0 | 6.2 | 1173 | 14.7 | 14.4 |
| LYSO Tefl. wrap. | 20.0 | 7.0 | 749 | 9.4 | 9.3 |

- Table II

##### Figure captions

Fig. 1. Schematic illustration of the 3D axial HPD-PET concept. Arrays of axially arranged long crystals are read out on both sides by HPDs.

Fig. 2. Schematic set-up used to measure effective light attenuation lengths.

Fig. 3. Energy spectra for polished YAP (upper panel) and LYSO (lower panel) scintillators exposed to 511 keV -rays impinging at the centre of the crystal lateral surface. The spectra have been obtained by summing the crystal left and right ADC signals in coincidence with the BaF2 signal. The shown spectra have been obtained with same high voltages for the left and for the right H3164-10 PMTs.

Fig. 4. Bidimensional spectrum (left versus right ADC channels) for a 3.2×3.2×100 mm3 polished YAP crystal exposed to 511 keV -rays at the centre (z = 5 cm) of its lateral surface. The spectrum has been taken in coincidence with a BaF2 crystal detecting the second 511 keV -ray from a  22Na source.

Fig. 5. The 511 keV -ray photopeak signal amplitude (log scale) measured at one end of the bar (here left) versus the z position of the 22Na source on the lateral surface of a polished YAP and LYSO crystal. The different slope of the two sets of points is due to the different eff value for the two scintillators.

Fig. 6. Pulse height of the photopeak centroids as delivered by a H3164-10 PMT for 511 keV -rays impinging laterally on a 10 cm long YAP crystal at different z positions. Full squares refer to a polished crystal, full points to a Teflon wrapped crystal, upward full triangles to a crystal coated with 1 nm Cr, downward full triangles to a white painted crystal, upward empty triangles to a crystal coated with 1.5 nm Au, empty squares to a black painted crystal, empty circles to a crystal coated with 3 nm Cr. The lines on the experimental points are exponential fits to the data, giving the light attenuation lengths eff reported in Table I. The extrapolation of the fits to z = 0 gives a rough estimate of N0/2 (see Eq. 1). The N0 parameter is obviously also affected by the surface properties of the crystal.

Fig. 7. The experimental eff values (points) obtained for polished (point at tCr = 0 nm) and Cr coated YAP crystals compared with the prediction (full line) of a photo tracking code [9].

Fig. 8. The reconstructed z positions (points) versus the real ones,  for a polished YAP (upper panel) and LYSO (lower panel) crystal. The reconstruction is restricted to photopeak events.  The continuous lines are linear fits to the data.

Fig. 9.  The distributions of the reconstructed z points for individual 511 keV -rays in a Teflon wrapped YAP crystal at three distances z = 2, 5, and 8 cm for top, central, and lower panel, respectively.

Fig. 10.  Uncertainty of z reconstruction (standard deviations of the z-reconstructed distributions) versus the source z position for YAP (left side) and LYSO (right side), polished (top panels) and Teflon wrapped (bottom panels) crystals. Same abscissa scales in left and right panels, different in top and bottom panels.

Fig. 11. The E/E energy resolutions(top panel), the z position resolutions (central panel), and the t time resolutions (bottom panel) measured with 511 keV -rays in the crystal centre (z = 5 cm) for YAP crystals at different eff values. All the resolutions have been evaluated from the standard deviations of the photopeaks in the sum spectra (left plus right).

Fig. 12. The E/E energy resolutions(top panel), the z position resolutions (central panel), and the t time resolutions (bottom panel) measured with 511 keV -rays in the crystal centre (z = 5 cm) for LYSO crystals at different eff values. All the resolutions have been evaluated from the standard deviation of the photopeak in the sum spectra (left plus right).

Fig. 1

Fig. 2

Fig. 3

Fig. 4

Fig. 5

Fig. 6

Fig. 7

YAP

LYSO

Fig. 8

Fig. 9

Fig. 10

Fig. 11

Fig. 12

Corresponding author, e-mail: deleo@ba.infn.it, fax number = +39 080 553 4938

Hamamatsu Photonics K.K., Japan.